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Tender reference: 2026-064
[bookmark: _Toc459799307]Annex 2 – Bid Form completed signed and stamped
· All bids must be submitted in typed format. Handwritten bids will not be accepted.
· The contract will be awarded to the lowest-priced bid that is technically compliant and meets all tender requirements.
· The bidder is not permitted to alter the section labelled “LRC to complete” below. Bidder is not allowed to use any other template.  
· In the event of any discrepancy between the unit price and the total price, LRC reserves the right to rely on the unit price stated in the financial bid and recalculate the total price accordingly, without requesting clarification from the bidder.
· The unit price shall include all costs related to labor, materials, equipment, packaging, transportation, delivery to LRC warehouse, installation, testing, commissioning, and all applicable bank transfer charges.
	LRC To Complete
	Bidder to Complete 

	Item # 
	Item/Milestone Required
	Unit of Measurement (UOM)
	Estimated 
Quantity 

	Unit Price in USD
Exclusive of VAT
	VAT (%)
	Unit Price in USD
 Inclusive of VAT
	Total Price in USD
Inclusive of VAT
	Lead time of delivery 
(in calendar days) from the date of each Purchase Order.

	1
	Conversion, Outfitting, Testing, and Commissioning of LRC-Provided IVECO Daily Van 50C15V (18m³) into a Fully Equipped Mobile Medical Clinic
	Van Conversion Units
	3
	
	
	
	
	








	




Annex 3: Technical Specifications:
	
Subject: Conversion, Outfitting, Testing, and Commissioning of LRC-Provided IVECO Daily Van 50C15V (18m³) into a Fully Equipped Mobile Medical Clinic


	Vehicle: Vehicle: IVECO Daily Van 50C15V – 18m³, Model Year 2026, – Euro 3 – Twin Rear Wheels 

	Requirements

	The Mobile Clinic is divided into three (3) sections:
1- Driver’s section
2- Reception area
3- Consultation/Examination area

	1. Driver’s section:

	No changes shall be made on this section. It shall be separated by a partition, contains driver’s and passenger’s seats and a 2Kg fire extinguisher

	2. Reception area:

	This area shall be used for patient reception and quick treatments.
It shall be located behind the driver’s section and separated by a properly fixed partition. 
The partition shall be made of 15mm plywood covered with 1mm formaica on both sides.

	The reception area shall contain the following:
· Three (3) flip seats fixed in the reception compartment. Seats shall have fixed backrest, foldable cushion seat, moving armrest, adjustable headrest, 3-point safety seatbelt, floor mounting, robust mechanical structure, easy to clean and UV-protected high-quality upholstery.
· Foldable wooden table shall be made of 15mm plywood covered with 1mm formaica on both sides.
· Overhead mounted cabinets shall be made of 15mm plywood covered with 1mm formaica on both sides with plexi-glass windows.
· Cabinet shall be made of 15mm plywood covered with 1mm formaica on both sides with at least 4 drawers having high-quality locking system.
· 2 LED roof bright lights.
· Electric outlets: 2x 12VDC and 2x 220VAC.
· Air extrication fan fixed to the left side of the vehicle, shall have a stainless-steel hood with a proper rubber sealing to prevent corrosion, water dissipation and air blockage.

	3. Examination area:

	The examination area shall be located directly behind the reception area. It shall be separated by a wooden partition made of 15mm plywood covered with 1mm formaica on both sides and an integrated door. Used for patient examination and treatments.

	The examination area shall contain the following:
· Examination bed with top made of special high quality foam material covered with high quality leather upholstery.
· Cabinets and drawers under the examination bed made of 15mm plywood covered with 1mm formaica on both sides can be used for storage of medical and non-medical equipment or material.
· Overhead console along the examination bed containing: 3x 220VAC outlets, 3x 12VDC outlets, 2x oxygen outlets with 2x flowmeters and humidifiers shall be connected to the central oxygen system, fixed suction unit with 1 liter Jar.
· 4 LED roof bright lights.
· Handlebar made round tube of minimum diameter 25mm and minimum thickness 2mm fixed on the inner roof by counter plates.
· Air extrication fan.
· Water sink made of stainless-steel grade 316 and minimum thickness 1.2mm.
· Large wooden cabinet made of 15mm plywood covered with 1mm formaica on both sides can be used for storage of medical and non-medical equipment or material, fixed to the side of the vehicle.
· Sharp container.
· Fire extinguisher 2kg.
Refrigerator
· Size (Cubic Feet): 5 Cft
· Capacity (Litres): 110 L
· Type: 1 Door
· Compressor Type: Universal
· Freestanding / Built-In: Freestanding
· Cooling Type: Defrost
· Power: 220VAC
· Color: White 
· Dimensions (W x H x D): 46 x 86 x 48
· Approximate Weight (KG): 20Kg

	General Requirements:

	· Floor of the entire mobile clinic shall be made of marine plywood of minimum 18mm thickness and at least 10 plies, covered with a washable and an anti-slip vinyl material of minimum thickness 1mm, shall have a matching color as the vehicles’ interior.
· All inner walls and inner roof of the vehicle to be covered with panel sheets (ABS or GRP) including thermal and sound isolation underneath.
· Inverter/charger of minimum 2000W power with a power saver and full load modes to ensure safety of the electric appliances.
· 2x batteries deep cycle batteries of minimum capacity 100Ah each.
· Complete electric dashboard including all necessary circuit breakers, fuses and electric connections (European brand) to control all the electric equipment installed in the mobile clinic.
· Auto-Eject automatic shoreline power connector equivalent, suitable for emergency and specialist vehicles, rated for 230 VAC, 50 Hz, minimum 16 A capacity, IEC 60309 connector configuration, automatic plug ejection upon engine start, compatible with both 12 VDC and 24 VDC vehicle systems, weatherproof protective cover, , vehicle-grade vibration-resistant construction, complete with mating plug, wiring harness, mounting hardware, and all accessories required for installation and commissioning.
· Extension electric cable minimum 20m long.
· 2x Stationary rooftop air conditions specially made for automotive use, inverter compressor technology, minimum cooling capacity 3,000 W and heating capacity 2,500 W, operating on 220–240 VAC, 50/60 Hz, maximum current consumption not exceeding 7 A in cooling mode, electronic thermostat, soft-start function, R32 refrigerant, airflow not less than 480 m³/h, operating temperature range from -2°C to +52°C, low-noise operation, complete with air distribution box, controls, mounting kit, electrical connections, and all accessories required for installation and commissioning in a mobile medical clinic vehicle.
· LED emergency warning lightbar, having an ultra-low-profile design, minimum length 1200 mm (48 inch), operating on 12 VDC, incorporating multi-color LED technology, minimum 28 flash patterns, front and rear flood lighting capability, alley lights, low-power mode, operating temperature range -40°C to +65°C, IP69K environmental protection, compliant with SAE J595 Class 1 and SAE J845 Class 1 standards, with mounting hardware and wiring harness.
· Manual metal sidestep to be installed underneath the sliding door.
· Rear view camera with a screen.
· Rear parking sensors.

	Additional Requirements

	Retractable box-type shadow awning, vehicle-mounted construction, minimum dimensions 4.0 m width × 3.0 m projection, aluminium cassette housing with articulated folding arms, UV-resistant waterproof fabric, corrosion-resistant powder-coated finish, wind resistance Class 2 or higher according to EN 13561, complete with mounting brackets, electrical controls, wiring harness, and all accessories required for installation on a mobile medical clinic vehicle.

The contractor shall be responsible for the complete conversion, installation, integration, testing, and commissioning of the mobile medical clinic. The final vehicle shall be delivered fully operational and ready for use. Any materials, components, accessories, or works required for proper operation and compliance with these specifications, whether explicitly mentioned or not, shall be considered included in the bidder’s offer.

	Technical Offer Requirements
The bidder shall submit a technical offer including, as a minimum:
· Detailed technical compliance statement against all requirements of Annex 3.
· Proposed mobile clinic layout/drawing showing the arrangement of all sections, equipment, and installations.
· Description of the conversion methodology, materials, and equipment to be used.
· Technical datasheets/catalogues and manufacturer details for major installed components.
· Proposed delivery schedule and conversion lead time.
· Warranty details and after-sales support.
Any deviation from the technical specifications shall be clearly identified and explained. Failure to provide the required technical information may result in rejection of the offer.
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