## Annex 2 – Bid Form completed signed and stamped

- All bids must be typed; handwritten bids will not be considered.

- The contract will be awarded to the lowest-cost technically compliant complete lot/items

- The quantities mentioned below are required unless there is a variation within a specified range with proper justification. The LRC reserves the right to increase or decrease the quantities.

- The bidder is not permitted to alter the section labelled “LRC to complete” below. Bidder is not allowed to use any other template.

- In the event of any discrepancy between the unit price and the total price, LRC reserves the right to rely on the unit price as stated in the financial bid, rather than the total price without go back to the supplier. The unit price must be provided with no more than two decimal places.

- Only one option is permitted per item. If a bidder submits more than one option, LRC will consider only the first option submitted For technical and Financial Evaluation.

- LABELED AND DATASHEETS FOR EACH ITEM MUST BE SUBMITTED

Lot 1: Medical Equipment

| LRC to Complete (BIDDERS ARE NOTE ALLOWED TO MODIFY THE LRC SECTION) |  |  |  | Bidder To Complete |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Item Number | Item/Milestone Required | Unit of Measurements | Estimated Qty | Unit Price in USD, Exclusive VAT | VAT (%) | Unit Price in USD, inclusive VAT | Total Price in USD, inclusive VAT | Lead time of Delivery for any Quantity from the date of the purchase order |
| 1 | Autoclave/sterilizer | Piece | 7 |  |  |  |  |  |
| 2 | Axillary crutches - Large | Pair | 400 |  |  |  |  |  |
| 3 | Axillary crutches - medium | Pair | 500 |  |  |  |  |  |
| 4 | Carts (multipurpose) | piece | 25 |  |  |  |  |  |
| 5 | Cleaning trolley | Piece | 7 |  |  |  |  |  |
| 6 | digital weight scale (adult) | Piece | 15 |  |  |  |  |  |
| 7 | Disinfectant soaking container | Piece | 40 |  |  |  |  |  |
| 8 | Dissecting | Piece | 0 |  |  |  |  |  |
| 9 | Doppler (portable) | Piece | 36 |  |  |  |  |  |
| 10 | Dressing Forceps | Piece | 14 |  |  |  |  |  |
| 11 | EKG | Piece | 7 |  |  |  |  |  |
| 12 | Elbow crutches mobile - Large | Piece | 200 |  |  |  |  |  |
| 13 | Elbow crutches mobile - Medium | Piece | 200 |  |  |  |  |  |
| 14 | Exam table | Piece | 28 |  |  |  |  |  |
| 15 | Exam table (OBS-GYN) | Piece | 7 |  |  |  |  |  |
| 16 | Examination Table – Foldable | Piece | 7 |  |  |  |  |  |
| 17 | Fetal Doppler (portable) | Piece | 10 |  |  |  |  |  |
| 18 | Folding Commode Toilet chair | Piece | 300 |  |  |  |  |  |
| 19 | Height board (pediatric) | Piece | 7 |  |  |  |  |  |
| 20 | Hemocue portable | Piece | 16 |  |  |  |  |  |
| 21 | Hemoglobin meter | Piece | 16 |  |  |  |  |  |
| 22 | IUD insertion kit: Ringed Forceps | Piece | 100 |  |  |  |  |  |
| 23 | IUD insertion kit: Scissors | Piece | 100 |  |  |  |  |  |
| 24 | IUD insertion kit: Tenaculum forceps straight | Piece | 100 |  |  |  |  |  |
| 25 | IUD insertion kit: Hysterometer | Piece | 100 |  |  |  |  |  |
| 26 | Kelly-Rankin Forceps | Piece | 14 |  |  |  |  |  |
| 27 | Kidney dish | Piece | 28 |  |  |  |  |  |
| 28 | Kocher artery forceps | Piece | 14 |  |  |  |  |  |
| 29 | Light pen | Piece | 20 |  |  |  |  |  |
| 30 | Measuring tape | Piece | 20 |  |  |  |  |  |
| 31 | Measurment scale (adult) | Piece | 7 |  |  |  |  |  |
| 32 | Tablet Counting Triangle | piece | 50 |  |  |  |  |  |
| 33 | Medication tablets counter / pills counter tray | Piece | 50 |  |  |  |  |  |
| 34 | Mobile exam light | Piece | 7 |  |  |  |  |  |
| 35 | Opthalmoscope | Piece | 24 |  |  |  |  |  |
| 36 | Otoscope | Piece | 7 |  |  |  |  |  |
| 37 | Oximeter | Piece | 7 |  |  |  |  |  |
| 38 | Quadripode cane | Piece | 400 |  |  |  |  |  |
| 39 | Reflex hammer | Piece | 7 |  |  |  |  |  |
| 40 | Scale Baby Digital 20 Kgs | Piece | 7 |  |  |  |  |  |
| 41 | Sphygmomanometer | Piece | 7 |  |  |  |  |  |
| 42 | Spot Vital Signs Device with mobile stand | Piece | 15 |  |  |  |  |  |
| 43 | Step stool | Piece | 35 |  |  |  |  |  |
| 44 | Stethoscope | Piece | 7 |  |  |  |  |  |
| 45 | Stick | Piece | 300 |  |  |  |  |  |
| 46 | Thermometers | Piece | 20 |  |  |  |  |  |
| 47 | Trays | Piece | 28 |  |  |  |  |  |
| 48 | Tripod | Piece | 300 |  |  |  |  |  |
| 49 | Tuning forks | Piece | 36 |  |  |  |  |  |
| 50 | Ultrasound (cardiac, abdominal, vaginal) | Piece | 7 |  |  |  |  |  |
| 51 | Waste Refrigerator | Piece | 7 |  |  |  |  |  |
| 52 | Wheel Chair - pneumatic tires | Piece | 500 |  |  |  |  |  |
| 53 | Wheel Chair - Solid tires | Piece | 600 |  |  |  |  |  |

## ANNEX 3: DETAILED SPECIFICATIONS

| ITEM # | Item Description | Item/Milestone Required |
| --- | --- | --- |
| 1 | Autoclave/sterilizer | Vacuum steam sterilizer. Bench mounted model. Front loading. High-grade stainless steel inner chamber. Chamber volume approx. 23 L. Cycle programs for Wrapped & Unwrapped Water Storage Tank: Sterilizer shall have inbuilt water reservoir with storage capacity up to 5 liters. The water reservoir shall have easy access for cleaning & to avoid bio film. Temperature range 105°C -135°C. Timer: 0-60 min. Safety system (over pressure, overheating, opening door, low water...). Automatic power Cut-off safety system. Automatic drying cycle. Digital or LCD display (Temperature, timer, Program Cycle). Led indicator (water, Heat, Sterilization, dry, end cycle). Integrated data printer. USB port The sterilizer unit shall come with wire mesh basket to hold the materials & suitable size instrument trays. Power supply: 220-240V 50 Hz. FDA or CE Approved |
| 2 | Axillary crutches - Large | Aluminum lightweight - Height adjustable from 175 - 195 cm approx. (Large) - High density rubber horizontal bar cushioned fitting for the underarm - Adjustable handgrip minimum 3 holes with a soft cover rounded rubber - Rubber ferrule replaceable - ISO certificate |
| 3 | Axillary crutches - medium | Aluminum lightweight - Height adjustable from 155-175 cm approx. (Medium) - High density rubber horizontal bar cushioned fitting for the underarm - Adjustable handgrip minimum 3 holes with a soft cover rounded rubber - Rubber ferrule replaceable - ISO certificate |
| 4 | Carts (multipurpose) | Stainless steel guardrails around. Minimum 1 drawer and preferable with guardrails around. Minimum 1 lower shelf with 4 wheels and brakes Length/Width: 60-90 cm Depth 45-60cm Height 80–95 cm |
| 5 | Cleaning trolley | Heavy duty plastic 2 buckets system for mops with a press, 2 small buckets for cloths 3 shelves Caution wet floor 2 Waste bags (Regular and infectious) Mop holder strong hook Length 100–115 cm, width 50–60 cm (maximum 60 cm), height 100–110 cm |
| 6 | digital weight scale (adult) | Adult Digital Floor Scale: Portable digital weighing scale for adults, displaying weight in kilograms (kg), with a minimum capacity of 200 kg . Equipped with a large LCD display, non-slip platform, automatic zero setting, battery operation, and a lightweight design suitable for Mobile Medical Unit (MMU) use. Compliant with Medical Device Directive 93/42/EEC or equivalent standards |
| 7 | Disinfectant soaking container | Non autoclavable with basket, transparent lid, durable plastic material (L × W × H): 35–40cm × 15-20cm × 10-12 cm |
| 8 | Dissecting | Anatomy forceps, dressing, standard, 14 cm, straight,Stainless steel, 400 grade,Passivated, maker mark. CE/FDA |
| 9 | Doppler (portable) | Portable / Handheld equipment for vascular assessment, and detection of peripheral arterial disease. Provide high quality blood flow information. Durable Long-Lasting Sealed pencil shaped vascular probes with frequencies selected from 5 MHz. Provide access to both deep and superficial arteries and veins for determining systolic blood pressures, vessel patency and locating hard-to-find vessels. Built-in High-fidelity speaker and auxiliary output for headphone. Integrated probe holder, carry case. Integrated gain control and Volume control. |
| 10 | Dressing Forceps | Stainless steel. Flat Handle With Spring Action. Straight serrated jaws. Plain - 16 cm CE/ISO |
| 11 | EKG | 12-lead ECG, simultaneous acquisition of all 12 leads, minimum 3-channel display and print Suitable for adult and pediatric patients Built-in color screen, minimum 5 inches AC mains operation (220–240 V, 50 Hz) Internal rechargeable lithium battery, minimum 100 ECGs or 3 hours of operation per charge Battery level indicator Automatic ECG measurement and interpretation (adult and pediatric) Manual and automatic recording modes Filters: AC (50 Hz), muscle, and baseline drift Internal memory for minimum 100 ECG records Patient data entry (name, ID, age, sex) Built-in thermal printer, paper width minimum 80 mm 1 patient cable, 10-lead 1 set of reusable limb clamps (4) and chest suction electrodes (6) power cable Additional option for A4 printing is preferable preferable to include clip-type (snap/tab) patient cable for single use patches |
| 12 | Elbow crutches mobile - Large | Size: 93cm/116cm -single - adjustable- with PVC handles and cuff - lightweight aluminum - closed mobile cuff -Angled neck- contoured hand grip - anti-slip rubber tip |
| 13 | Elbow crutches mobile - Medium | Size: 87cm/110cm - single- adjustable- with PVC handles and cuff - lightweight aluminum - closed mobile cuff - angled neck- contoured hand grip - anti-slip rubber tip |
| 14 | Exam table | Approx. Dimensions: 180 cm l x 60cm w x 80cm h - Frame: Stainless steel tubesØ 30mm. - Adjustable backrest by hand lever - Two sectioned-simulated leather, Mattress, fixed, without seams, 40mm thick. Equipped with paper roll holder. Manual. CE mark |
| 15 | Exam table (OBS-GYN) | Approx. Dimensions: 180 cm l x 60cm w x 80cm h - Frame: Stainless steel tubesØ 30mm. - Adjustable backrest by hand lever - Three sections-simulated leather, Mattress, fixed, without seams, 40mm thick. Removable Legrest. Equipped with paper roll holder. Manual. CE mark |
| 16 | Examination Table – Foldable | Portable foldable examination table suitable for mobile medical units. Lightweight aluminum or steel frame with padded foam mattress covered with waterproof, easy-to-clean medical-grade vinyl. Foldable design with carrying handle for easy transport and storage. Stable, durable, and capable of supporting a minimum load of 150 kg. Dimensions approximately 180–190 cm (L) × 60–70 cm (W) × 70–80 cm (H). CE marked |
| 17 | Fetal Doppler (portable) | Portable fetal heart beat detector, lightweight , Frequency : 2.0Mhz/2.5 Mhz , Intensity: 5m W/cm or less. FDA or CE mark. |
| 18 | Folding Commode Toilet chair | Foldable plated frame lightweight with backrest. Height adjustable with push pin minimum 6 holes. Four swivel casters, minimum 4" (100 mm) diameter with individual brakes. Armrests with a secure grip. Pail with carry handle and cover. Padded PVC seat with cover. Base width 110cm. ISO/FDA |
| 19 | Height board (pediatric) | Portable height measuring baby scale, Range:10-80 cm. Graduation: 1mm. Fixed head piece (headboard). Slide foot positioner. Washable |
| 20 | Hemocue portable | Precise factory calibration against the ICSH reference method. Blood based liquid control. Built in self test. Capillary, venous, or arterial whole blood sample. Dimensions 160 X 43 X 85 mm. With batteries installed. Modified azide methemoglobin reaction dual wavelenghts 570nm and 880nm for compensation of turbidity. Measurement Range: 0–25.6 g/dL (0–256 g/L, 0–15.9 mmol/L). Result within 60 seconds. Sample Volume: ~10 µL. Materials included; Hemocue + analyzer, Hemocue cleaner, AC adapter, User manual. ISO certificate. |
| 21 | Hemoglobin meter | Automation: Fully Automatic Number Of Chambers 1 Sample Mode: Whole Blood Method: Flow Cytometry Chamber Type: Single Portable, able to work in hot climate, up to 45 degree Celsius. Easy To Use FDA or CE mark. |
| 22 | IUD insertion kit: Ringed Forceps | 24cm, serrated jaws, straight, Stainless steel, 400 grade, Passivated, maker mark, CE/FDA |
| 23 | IUD insertion kit: Scissors | Scissors Mayo, straight, with blunt-tipped blades. Length: 15 cm. Stainless steel, 400 grade,Passivated, maker mark, CE/FDA |
| 24 | IUD insertion kit: Tenaculum forceps straight | FORCEPS, SEIZING, POZZI, 25 cm, tenaculum, straight, Stainless steel, 400 grade, Passivated, maker mark, CE/FDA. Screw lock. Ratchet with variable settings. Adjustable grip. |
| 25 | IUD insertion kit: Hysterometer | 32 cm, sims curved, rigid, malleable, graduated in cm, sainless steel 400 grade, Passivated, maker mark, CE/FDA |
| 26 | Kelly-Rankin Forceps | Stainless steel. Straight lengths 16 cm. Serrated tips. Handle:Wider finger rings. CE/ISO |
| 27 | Kidney dish | Medium, length 200-2250 mm Width 90–110 mm Depth 40-45 mm Thickc 0.8–1.0 mm 400-500 ml capacity |
| 28 | Kocher artery forceps | Stainless steel. Ratcheted locking mechanism controls hemostat's jaw pressure. Straight lengths 16 cm. Serrated heavy tips. 1 X 2 teeth. CE/ISO |
| 29 | Light pen | Battery Powered. Light Source Type. LED. Material: Aluminum. Battery Cell Composition: Alkaline. |
| 30 | Measuring tape | Tailor type, PVC coated polyester, 20mm x 3m |
| 31 | Measurment scale (adult) | Adult Mechanical Floor scale Displays weight in kg, Capacity up to 200 kg.Graduation Weight: 100 g.Reading on both sides. Robust lever system.Height adjustable rod up.Height Rod Range: approx. 60 - 200 cm with 1 mm graduations. Anti-slip platform. Adjustable zero point.Weighing units (kg ). Transport castors-Portable |
| 32 | Tablet Counting Triangle | Material: ABS, acrylic, or stainless steel; anti-static; alcohol-resistant Size: Sides about 20–25 cm; raised edge on two sides Capacity: Minimum 13 rows (91 tablets) Markings: Permanent row-count markings Design: Seamless, easy to clean, non-slip base |
| 33 | Medication tablets counter / pills counter tray | Rigid ABS or acrylic, antimicrobial or anti-static Better to have spatula included side Channel with Hinged cover and a pouring spout (L × W × H): 24-30 mm x13-18mm x3-3.5 mm |
| 34 | Mobile exam light | Steel structure with epoxy coating - Source of light oriented to limitless positions - Voltage 220v - Cable lenght 150cm - White PVC base of 65cm with five rolling wheels of 40mm, two or three of them with brakes - LED light 10w 24º - Cold light - Ilumination 15000 Lux to 50cm - Colour rendering RA 80% - Colour temperature 3000ºK - Dimensions 65x65x125cm. ISO/FDA |
| 35 | Opthalmoscope | Direct opthalmoscope, battery powered, At least 2.5 V Xenon or Halogen light source. Magnification up to x15 . Anti-reflection lens. At least 3 apertures and fixation star. Dust free sealed optics and aspherical optical system. Red-free, blue and polarization filters. Supplied in protective, reclosable container. On/off switch to be robust and easy to use. External material to be non-ferrous. |
| 36 | Otoscope | Fiber Optic. Onyx w/ Handle complete, with 4 reusable specula different sizes: 2.5, 3, 4 and 5mm for adult and pediatric. Battery Type: AA Alkaline Magnification Power: 3.4X Hard case. Accessories: Set of plastic specula, varying diameters between 2.0 and 5.0 mm Two spare bulbs. At least 10 reusable (autoclavable) otoscope specula for each one of the following measure: 2, 3 and 5 mm. FDA or CE mark. |
| 37 | Oximeter | Portable, battery-powered SpO2 measurement range at least 70 to 99 %, minimum resolution 1% Accuracy of SpO2 better than ± 2% Pulse rate range at least 30 to 240 bpm, Accuracy of pulse rate better than ± 4 bpm Signal strength or quality to be visually displayed Automatic power-off facility required after minimum of 1 minute Low battery display required Fingertip model FDA, CE approved |
| 38 | Quadripode cane | Chrome plated steel - Lightweight approx. 600gr -Height adjustable - 690 to 940mm - Anodized aluminum neck with a molded plastic grip handle - Pyramid shape with base with four footprint approx. 165 X140mm with rubber approx. 3cm on the end -Suitable for left and right handed users. ISO certificate |
| 39 | Reflex hammer | Solid handle and a rubber head.Length: 20cm. |
| 40 | Scale Baby Digital 20 Kgs | Tray dimensions are 20.6 W x 10.5 D x 4 H inch (524 W x 267 D x 102 H mm). Display options include date and time as well as storage of previous weight on a large 1 inch (25 mm) high LCD screen. 44 lb x 0.5 oz / 20 kg x 10 g capacity. ISO/FDA |
| 41 | Sphyngomanometer | Aneroid pressure gauge (no stop pin) with an air pressure release-valve. Rubber bulb with palm fitting handle. Connector to rubber tube. Connecting tube.Tube length to be greater than 30cm Cuff surround to be removable and Washable with inflatable bladder. Included: Cuffs: (Adult, Child, obese ). FDA, CE approved |
| 42 | Spot Vital Signs Device with mobile stand | Portable, non-invasive spot vital signs machine. It shows completes BP, pulse rate, temperature, and SpO2 in approximately 30 seconds Portable - carry it, wheel it Big, bright, easy-to-read LCD screen Materials Included Spot Vital Signs Unit (1) Adult Cuff (1) Pediatric cuff (1) Large Adult Cuff (1) Blood Pressure Hose (1) Lithium Ion Battery (1) Power Cord SpO2 finger tip sensor Thermometer (oral/axillary/rectal) User Manual Technical Specifications AC Power Supply: 7.2 V DC at 1.0 A rechargeable sealed lead acid internal battery Battery Recharge Time: 90% to 100% after 12 hours Warranty Information Mobile stand for medical monitor use in PHC/clinical settings. Compatible with standard patient monitors (universal mounting plate or equivalent). Height adjustable (approx. 120–180 cm). Stable base with breaks Strong metal frame (steel or aluminum), rust-resistant. Minimum load capacity: 10–15 kg or suitable for monitor weight. better to be Equipped with cable management system. Easy to clean and disinfect. Stable and anti-tilt design for safe movement. |
| 43 | Step stool | 2 steps stool. Weight capacity: 150 kgs. Chrome-plated 18 gauge steel for durability. Skid-resistant rubber mat |
| 44 | Stethoscope | Dual head, Head and earpiece constructed of non-ferrous material Earpieces to have shaped, cushioned ends Tubes to be synthetic material |
| 45 | Stick | Lightweight aluminum - adjustment range between 75-100cm - PVC handle - anti-slip rubber tip |
| 46 | Thermometers | Measurement Sites: Oral, axillary, and rectal. Measurement Time: Approximately 4–6 seconds (oral), 10–13 seconds (rectal/pediatric axillary), 12–15 seconds (adult axillary). Accuracy: ±0.1 °C (±0.2 °F). Display: Large, backlit digital LCD with mode indicators and memory recall of last reading. Probes: Interchangeable color-coded probes (oral/axillary and rectal) with removable probe wells to minimize cross-contamination Probe Covers: Single-use, disposable, with push-button ejection system. Battery: Operates on 3 × AA batteries with capacity of approximately 6,000 measurements or rechargeble Dimensions/Weight: Approx. 215 mm × 81 mm × 62 mm; ~357 g. Operating Range: 10 °C to 40 °C ambient. Mounting Options: Handheld |
| 47 | Trays | Rectangular tray with rounded corners. Stainless steel. Smooth surface. 30cm*15-22 cm*3 cm. Steel quality, as per ISO standard |
| 48 | Tripod | Chrome plated steel - Lightweight approx. 600gr -Height adjustable - 690 to 940mm - Anodized aluminum neck with a molded plastic grip handle - Small base with three footprint approx. 165 X140mm with rubber approx. 3cm on the end -Suitable for left and right handed users ISO certificate |
| 49 | Tuning forks | Nickel-plated steel, material thickness 8 mm, for otologists, high-grade steel, nickle-plated approx. 17 cm |
| 50 | Ultrasound (cardiac, abdominal, vaginal) | Color Doppler Ultrasound - Display monitor: minimum 15" color high resolution LCD - Articulating monitor arm - Console design, full size keyboard, 8 TGC pods, Integrated speakers - Probe holders, Gel holder, front and rear handles - Minimum 3 probe connectors - Minimum Built-in storage (HDD) (500GB) - 2 USB ports, DVD-RW driver - Applications: Abdominal, Breast, Obstetrical/Gynecological, cardiac Real-time automatic doppler calculations 2D imaging - Automatic IMT - Probes: Convex transducer, Linear Array ,cardiac adult transducer, Transvaginal - Connectivity: Ethernet port, VGA port - B/W thermal Printer - Electrical power: Voltage 110-220 V AC Supplied on stable, mobile trolley fitted with 4 wheels that can be braked FDA or CE mark. |
| 51 | Waste Refrigerator | Cabinet Type: Plastic liner o Volume: 550(Liter) o Dimensions: (WxDxH) 75x76x180(cm) o Number of Shelves: Minimum 3 o No Freezer FEATURES: o Dynamic cooling system o Tropical compressor resist voltage fluctuation. o Multi positional shelf system o Environmentally friendly o Adjustable high load bearing shelves o Pre painted steel inner liner o Heavy duty unidirectional wheels o Adjustable feet in the front o Uniform air distribution o Wide steel condenser with cooling fan for high efficiency and durability o Clean back design o Non PVC door frame o 220/240V 50/60HZ o Durable ABS lower grill |
| 52 | Wheel Chair - pneumatic tires | Light Weight - Removable wheelchair arms - Soft padded washable upholstery - Detachable swing away wheelchair foot rest - Pneumatic wheelchair tires - Wheels 21’’ + (60cm) in diameter - Locking brakes on each wheel - Detachable wheelchair arms -Folding Back - Overall width: 64cm - Seat width: 45-55cm - Seat depth: 40cm - ISO certificate |
| 53 | Wheel Chair - Solid tires | Light Weight - Removable wheelchair arms - Soft padded washable upholstery - Detachable swing away wheelchair foot rest - Solid tires - Wheels 21’’ + (60cm) in diameter - Locking brakes on each wheel -Folding Back - Overall width: 64cm - Seat width: 45-55 cm - Seat depth: 40cm - ISO certificate |