National tender application pack V0.3Lebanese Red Cross 

Tender reference: 2026-093
[bookmark: _Toc459799307]Annex 2 – Bid Form completed signed and stamped
· All bids must be typed; handwritten bids will not be considered.
· The contract will be awarded to the lowest-cost technically compliant complete lot/items
· The quantities mentioned below are required unless there is a variation within a specified range with proper justification. The LRC reserves the right to increase or decrease the quantities.
· The bidder is not permitted to alter the section labelled “LRC to complete” below. Bidder is not allowed to use any other template.  
· In the event of any discrepancy between the unit price and the total price, LRC reserves the right to rely on the unit price as stated in the financial bid, rather than the total price without go back to the supplier. The unit price must be provided with no more than two decimal places.
· Only one option is permitted per item. If a bidder submits more than one option, LRC will consider only the first option submitted For technical and Financial Evaluation. 
· LABELED AND DATASHEETS FOR EACH ITEM MUST BE SUBMITTED 

Lot 1: Medical Equipment 
	LRC to Complete (BIDDERS ARE NOTE ALLOWED TO MODIFY THE LRC SECTION)
	Bidder To Complete

	Item Number
	Item/Milestone Required
	Unit of Measurements
	Estimated Qty
	Unit Price in USD, Exclusive VAT 
	VAT (%)
	Unit Price in USD, inclusive VAT 
	Total Price in USD, inclusive VAT 
	Lead time of Delivery for any Quantity from the date of the purchase order 

	1
	Autoclave/sterilizer 
	Piece
	7
	 
	 
	 
	 
	 

	2
	Axillary crutches - Large
	Pair
	400
	 
	 
	 
	 
	 

	3
	Axillary crutches - medium 
	Pair
	500
	 
	 
	 
	 
	 

	4
	Carts (multipurpose) 
	piece
	25
	 
	 
	 
	 
	 

	5
	Cleaning trolley
	Piece
	7
	 
	 
	 
	 
	 

	6
	digital weight scale (adult)
	Piece
	15
	 
	 
	 
	 
	 

	7
	Disinfectant soaking container
	Piece
	40
	 
	 
	 
	 
	 

	8
	Dissecting 
	Piece
	0
	 
	 
	 
	 
	 

	9
	Doppler (portable)
	Piece
	36
	 
	 
	 
	 
	 

	10
	Dressing Forceps
	Piece
	14
	 
	 
	 
	 
	 

	11
	EKG
	Piece
	7
	 
	 
	 
	 
	 

	12
	Elbow crutches mobile - Large
	Piece
	200
	 
	 
	 
	 
	 

	13
	Elbow crutches mobile - Medium
	Piece
	200
	 
	 
	 
	 
	 

	14
	Exam table 
	Piece
	28
	 
	 
	 
	 
	 

	15
	Exam table (OBS-GYN)
	Piece
	7
	 
	 
	 
	 
	 

	16
	Examination Table – Foldable
	Piece
	7
	 
	 
	 
	 
	 

	17
	Fetal Doppler (portable)
	Piece
	10
	 
	 
	 
	 
	 

	18
	Folding Commode Toilet chair
	Piece
	300
	 
	 
	 
	 
	 

	19
	Height board (pediatric)
	Piece
	7
	 
	 
	 
	 
	 

	20
	Hemocue portable
	Piece
	16
	 
	 
	 
	 
	 

	21
	Hemoglobin meter 
	Piece
	16
	 
	 
	 
	 
	 

	22
	IUD insertion kit: 
Ringed Forceps 
	Piece
	100
	 
	 
	 
	 
	 

	23
	IUD insertion kit: 
Scissors
	Piece
	100
	 
	 
	 
	 
	 

	24
	IUD insertion kit: 
Tenaculum forceps straight 
	Piece
	100
	 
	 
	 
	 
	 

	25
	IUD insertion kit: Hysterometer
	Piece
	100
	 
	 
	 
	 
	 

	26
	Kelly-Rankin Forceps
	Piece
	14
	 
	 
	 
	 
	 

	27
	Kidney dish
	Piece
	28
	 
	 
	 
	 
	 

	28
	Kocher artery forceps
	Piece
	14
	 
	 
	 
	 
	 

	29
	Light pen
	Piece
	20
	 
	 
	 
	 
	 

	30
	Measuring tape 
	Piece
	20
	 
	 
	 
	 
	 

	31
	Measurment scale (adult)
	Piece
	7
	 
	 
	 
	 
	 

	32
	Tablet Counting Triangle 
	piece
	50
	 
	 
	 
	 
	 

	33
	Medication tablets counter / pills counter tray
	Piece
	50
	 
	 
	 
	 
	 

	34
	Mobile exam light
	Piece
	7
	 
	 
	 
	 
	 

	35
	Opthalmoscope
	Piece
	24
	 
	 
	 
	 
	 

	36
	Otoscope 
	Piece
	7
	 
	 
	 
	 
	 

	37
	Oximeter
	Piece
	7
	 
	 
	 
	 
	 

	38
	Quadripode cane
	Piece
	400
	 
	 
	 
	 
	 

	39
	Reflex hammer 
	Piece
	7
	 
	 
	 
	 
	 

	40
	Scale Baby Digital 20 Kgs 
	Piece
	7
	 
	 
	 
	 
	 

	41
	Sphygmomanometer 
	Piece
	7
	 
	 
	 
	 
	 

	42
	Spot Vital Signs Device with mobile stand
	Piece
	15
	 
	 
	 
	 
	 

	43
	Step stool
	Piece
	35
	 
	 
	 
	 
	 

	44
	Stethoscope
	Piece
	7
	 
	 
	 
	 
	 

	45
	Stick
	Piece
	300
	 
	 
	 
	 
	 

	46
	Thermometers 
	Piece
	20
	 
	 
	 
	 
	 

	47
	Trays
	Piece
	28
	 
	 
	 
	 
	 

	48
	Tripod
	Piece
	300
	 
	 
	 
	 
	 

	49
	Tuning forks
	Piece
	36
	 
	 
	 
	 
	 

	50
	Ultrasound (cardiac, abdominal, vaginal)
	Piece
	7
	 
	 
	 
	 
	 

	51
	Waste Refrigerator
	Piece
	7
	 
	 
	 
	 
	 

	52
	Wheel Chair - pneumatic tires
	Piece
	500
	 
	 
	 
	 
	 

	53
	Wheel Chair - Solid tires
	Piece
	600
	 
	 
	 
	 
	 








ANNEX 3: DETAILED SPECIFICATIONS 
	ITEM #
	Item Description
	Item/Milestone Required

	1
	Autoclave/sterilizer 
	Vacuum steam sterilizer. Bench mounted model. Front loading. 
High-grade stainless steel inner chamber. Chamber volume approx. 23 L.
Cycle programs for Wrapped & Unwrapped
Water Storage Tank: Sterilizer shall have inbuilt water reservoir with storage capacity up to 5 liters. The water reservoir shall have easy access for cleaning & to avoid bio film.
Temperature range 105°C -135°C. Timer: 0-60 min.
Safety system (over pressure, overheating, opening door, low water...).
Automatic power Cut-off safety system.
Automatic drying cycle.
Digital or LCD display (Temperature, timer, Program Cycle). Led indicator (water, Heat, Sterilization, dry, end cycle). Integrated data printer. USB port
The sterilizer unit shall come with wire mesh basket to hold the materials & suitable size instrument trays.
Power supply: 220-240V 50 Hz.
FDA or CE Approved

	2
	Axillary crutches - Large
	Aluminum lightweight - Height adjustable from 175 - 195  cm approx. (Large) - High density rubber horizontal bar cushioned fitting for the underarm  - Adjustable handgrip minimum 3 holes with a soft cover rounded rubber - Rubber ferrule replaceable - ISO certificate

	3
	Axillary crutches - medium 
	Aluminum lightweight - Height adjustable from 155-175 cm approx. (Medium) - High density rubber horizontal bar cushioned fitting for the underarm  - Adjustable handgrip minimum 3 holes with a soft cover rounded rubber - Rubber ferrule replaceable - ISO certificate

	4
	Carts (multipurpose) 
	Stainless steel  guardrails around.
Minimum 1 drawer  and preferable with guardrails around. 
Minimum 1 lower shelf
with 4 wheels and brakes
Length/Width: 60-90 cm 
Depth 45-60cm 
Height 80–95 cm

	5
	Cleaning trolley
	Heavy duty plastic
2 buckets system for mops with a press, 2 small buckets for cloths
3 shelves
Caution wet floor
2 Waste bags (Regular and infectious)
Mop holder strong hook
Length 100–115 cm, width 50–60 cm (maximum 60 cm), height 100–110 cm

	6
	digital weight scale (adult)
	Adult Digital Floor Scale: Portable digital weighing scale for adults, displaying weight in kilograms (kg), with a minimum capacity of 200 kg . Equipped with a large LCD display, non-slip platform, automatic zero setting, battery operation, and a lightweight design suitable for Mobile Medical Unit (MMU) use. Compliant with Medical Device Directive 93/42/EEC or equivalent standards

	7
	Disinfectant soaking container
	Non autoclavable with basket, transparent lid, durable plastic material
(L × W × H): 35–40cm × 15-20cm × 10-12 cm

	8
	Dissecting 
	Anatomy forceps, dressing, standard, 14 cm, straight,Stainless steel, 400 grade,Passivated, maker mark. CE/FDA

	9
	Doppler (portable)
	Portable / Handheld equipment for vascular assessment, and detection of peripheral arterial disease. Provide high quality blood flow information. Durable Long-Lasting Sealed pencil shaped vascular probes with frequencies selected from 5 MHz. Provide access to both deep and superficial arteries and veins for determining systolic blood pressures, vessel patency and locating hard-to-find vessels.
Built-in High-fidelity speaker and auxiliary output for headphone.
Integrated probe holder, carry case.
Integrated gain control and Volume control.

	10
	Dressing Forceps
	Stainless steel. Flat Handle With Spring Action. Straight serrated jaws.  Plain   - 16 cm CE/ISO

	11
	EKG
	12-lead ECG, simultaneous acquisition of all 12 leads, minimum 3-channel display and print
Suitable for adult and pediatric patients
Built-in color screen, minimum 5 inches
AC mains operation (220–240 V, 50 Hz)
Internal rechargeable lithium battery, minimum 100 ECGs or 3 hours of operation per charge
Battery level indicator
Automatic ECG measurement and interpretation (adult and pediatric)
Manual and automatic recording modes
Filters: AC (50 Hz), muscle, and baseline drift
Internal memory for minimum 100 ECG records
Patient data entry (name, ID, age, sex)
Built-in thermal printer, paper width minimum 80 mm
1 patient cable, 10-lead
1 set of reusable limb clamps (4) and chest suction electrodes (6)
power cable
Additional option for A4 printing is preferable
preferable to include clip-type (snap/tab) patient cable for single use patches

	12
	Elbow crutches mobile - Large
	Size:  93cm/116cm -single - adjustable- with PVC handles and cuff - lightweight aluminum - closed mobile cuff -Angled neck- contoured hand grip - anti-slip rubber tip

	13
	Elbow crutches mobile - Medium
	Size: 87cm/110cm - single- adjustable- with PVC handles and cuff - lightweight aluminum - closed mobile cuff -   angled neck- contoured hand grip - anti-slip rubber tip

	14
	Exam table 
	Approx. Dimensions: 180 cm l x 60cm w x 80cm h - Frame: Stainless steel tubesØ 30mm. - Adjustable backrest by hand lever - Two sectioned-simulated leather, Mattress, fixed, without seams, 40mm thick. Equipped with paper roll holder. Manual.
CE mark

	15
	Exam table (OBS-GYN)
	Approx. Dimensions: 180 cm l x 60cm w x 80cm h - Frame: Stainless steel tubesØ 30mm. - Adjustable backrest by hand lever - Three sections-simulated leather, Mattress, fixed, without seams, 40mm thick. Removable Legrest. Equipped with paper roll holder. Manual.
CE mark

	16
	Examination Table – Foldable
	Portable foldable examination table suitable for mobile medical units. Lightweight aluminum or steel frame with padded foam mattress covered with waterproof, easy-to-clean medical-grade vinyl. Foldable design with carrying handle for easy transport and storage. Stable, durable, and capable of supporting a minimum load of 150 kg. Dimensions approximately 180–190 cm (L) × 60–70 cm (W) × 70–80 cm (H). CE marked

	17
	Fetal Doppler (portable)
	Portable fetal heart beat detector, lightweight , Frequency : 2.0Mhz/2.5 Mhz , Intensity: 5m W/cm or less. FDA or CE mark.

	18
	Folding Commode Toilet chair
	Foldable plated frame lightweight with backrest. Height adjustable with push pin minimum 6 holes. Four swivel casters, minimum 4" (100 mm) diameter with individual brakes. Armrests with a secure grip. Pail with carry handle and cover. Padded PVC seat with cover. Base width 110cm. ISO/FDA                           

	19
	Height board (pediatric)
	Portable height measuring baby scale, Range:10-80 cm. Graduation:  1mm. Fixed head piece (headboard).
Slide foot positioner. Washable 

	20
	Hemocue portable
	Precise factory calibration against the ICSH reference method. Blood based liquid control. Built in self test. Capillary, venous, or arterial whole blood sample. Dimensions 160 X 43 X 85 mm. With batteries installed. Modified azide methemoglobin reaction dual wavelenghts 570nm and 880nm for compensation of turbidity. Measurement Range: 0–25.6 g/dL (0–256 g/L, 0–15.9 mmol/L). Result within 60 seconds. Sample Volume: ~10 µL. Materials included; Hemocue + analyzer, Hemocue cleaner, AC adapter, User manual. ISO certificate.

	21
	Hemoglobin meter 
	Automation: Fully Automatic
Number Of Chambers 1
Sample Mode: Whole Blood
Method: Flow Cytometry
Chamber Type: Single
Portable, able to work in hot climate, up to 45 degree Celsius.
Easy To Use
FDA or CE mark.

	22
	IUD insertion kit: 
Ringed Forceps 
	24cm, serrated jaws, straight, Stainless steel, 400 grade, Passivated, maker mark, CE/FDA

	23
	IUD insertion kit: 
Scissors
	Scissors Mayo, straight, with blunt-tipped blades. Length: 15 cm. Stainless steel, 400 grade,Passivated, maker mark, CE/FDA

	24
	IUD insertion kit: 
Tenaculum forceps straight 
	FORCEPS, SEIZING, POZZI, 25 cm, tenaculum, straight, Stainless steel, 400 grade, Passivated, maker mark, CE/FDA. Screw lock. Ratchet with variable settings. Adjustable grip.

	25
	IUD insertion kit: Hysterometer
	32 cm, sims curved, rigid, malleable, graduated in cm, sainless steel 400 grade,  Passivated,  maker mark, CE/FDA

	26
	Kelly-Rankin Forceps
	Stainless steel. Straight lengths 16 cm. Serrated tips. Handle:Wider finger rings. CE/ISO

	27
	Kidney dish
	Medium, 
length 200-2250 mm 
Width 90–110 mm 
Depth 40-45 mm
Thickc 0.8–1.0 mm
400-500 ml capacity 

	28
	Kocher artery forceps
	Stainless steel. Ratcheted locking mechanism controls hemostat's jaw pressure. Straight lengths 16 cm. Serrated heavy tips. 1 X 2 teeth. CE/ISO

	29
	Light pen
	Battery Powered.
Light Source Type. LED.
Material: Aluminum.
Battery Cell Composition: Alkaline.

	30
	Measuring tape 
	Tailor type, PVC coated polyester, 20mm x 3m

	31
	Measurment scale (adult)
	Adult Mechanical Floor scale
Displays weight in kg, Capacity up to 200 kg.Graduation Weight: 100 g.Reading on both sides.
Robust lever system.Height adjustable rod up.Height Rod Range: approx. 60 - 200 cm with 1 mm graduations.
Anti-slip platform.
Adjustable zero point.Weighing units (kg ).
Transport castors-Portable

	32
	Tablet Counting Triangle 
	Material: ABS, acrylic, or stainless steel; anti-static; alcohol-resistant
Size: Sides about 20–25 cm; raised edge on two sides
Capacity: Minimum 13 rows (91 tablets)
Markings: Permanent row-count markings
Design: Seamless, easy to clean, non-slip base

	33
	Medication tablets counter / pills counter tray
	Rigid ABS or acrylic,  antimicrobial or anti-static
Better to have spatula included
side Channel with Hinged cover and a pouring spout
(L × W × H): 24-30 mm x13-18mm x3-3.5 mm

	34
	Mobile exam light
	Steel structure with epoxy coating - Source of light oriented to limitless positions - Voltage 220v - Cable lenght 150cm - White PVC base of 65cm with five rolling wheels of 40mm,  two or three of them with brakes - LED light 10w 24º - Cold light - Ilumination 15000 Lux to 50cm - Colour rendering RA 80% - Colour temperature 3000ºK - Dimensions 65x65x125cm. ISO/FDA

	35
	Opthalmoscope
	Direct opthalmoscope, battery powered, 
At least 2.5 V Xenon or Halogen light source.
Magnification up to x15 .  
Anti-reflection lens.
At least 3 apertures and fixation star.
Dust free sealed optics and aspherical optical system.
Red-free, blue and polarization filters.
Supplied in protective, reclosable container.
On/off switch to be robust and easy to use.
External material to be non-ferrous.

	36
	Otoscope 
	Fiber Optic. Onyx w/ Handle complete, with 4 reusable specula different sizes: 2.5, 3, 4 and 5mm for adult and pediatric.  
Battery Type: AA Alkaline
Magnification Power: 3.4X
Hard case. Accessories: Set of plastic specula, varying diameters between 2.0 and 5.0 mm
Two spare bulbs. At least 10 reusable (autoclavable) otoscope specula for each one of the following measure: 2, 3 and 5 mm.
FDA or CE mark.

	37
	Oximeter
	Portable, battery-powered
SpO2 measurement range at least 70 to 99 %, minimum resolution 1%
Accuracy of SpO2 better than ± 2%
Pulse rate range at least 30 to 240 bpm, 
Accuracy of pulse rate better than ± 4 bpm
Signal strength or quality to be visually displayed
Automatic power-off facility required after minimum of 1 minute
Low battery display required
Fingertip model
FDA, CE  approved

	38
	Quadripode cane
	Chrome plated steel - Lightweight approx. 600gr -Height adjustable -  690 to 940mm - Anodized aluminum neck with a molded plastic grip handle - Pyramid shape with  base with four footprint approx. 165 X140mm with rubber approx. 3cm on the end -Suitable for left and right handed users. ISO certificate

	39
	Reflex hammer 
	Solid handle and a rubber head.Length: 20cm.

	40
	Scale Baby Digital 20 Kgs 
	Tray dimensions are 20.6 W x 10.5 D x 4 H inch (524 W x 267 D x 102 H mm). Display options include date and time as well as storage of previous weight on a large 1 inch (25 mm) high LCD screen. 44 lb x 0.5 oz / 20 kg x 10 g capacity. ISO/FDA

	41
	Sphyngomanometer 
	Aneroid pressure gauge (no stop pin) with an air pressure release-valve. Rubber bulb with palm fitting handle. Connector to rubber tube. Connecting tube.Tube length to be greater than 30cm
Cuff surround to be removable and Washable with inflatable bladder. Included: Cuffs: (Adult, Child, obese ). FDA, CE  approved

	42
	Spot Vital Signs Device with mobile stand
	Portable, non-invasive spot vital signs machine.  It shows completes BP, pulse rate, temperature, and SpO2 in approximately 30 seconds
Portable - carry it, wheel it
Big, bright, easy-to-read LCD screen
Materials Included
Spot Vital Signs Unit
(1) Adult Cuff
(1) Pediatric cuff
(1) Large Adult Cuff
(1) Blood Pressure Hose
(1) Lithium Ion Battery
(1) Power Cord
 SpO2 finger tip sensor
Thermometer (oral/axillary/rectal)
User Manual
Technical Specifications
AC Power Supply: 7.2 V DC at 1.0 A rechargeable sealed lead acid internal battery
Battery Recharge Time: 90% to 100% after 12 hours
Warranty Information
Mobile stand for medical monitor use in PHC/clinical settings.
Compatible with standard patient monitors (universal mounting plate or equivalent).
Height adjustable (approx. 120–180 cm).
Stable base with breaks
Strong metal frame (steel or aluminum), rust-resistant.
Minimum load capacity: 10–15 kg or suitable for monitor weight.
better to be Equipped with cable management system.
Easy to clean and disinfect.
Stable and anti-tilt design for safe movement.

	43
	Step stool
	2 steps stool. Weight capacity: 150 kgs. Chrome-plated 18 gauge steel for durability. Skid-resistant rubber mat

	44
	Stethoscope
	Dual head, Head and earpiece constructed of non-ferrous material
Earpieces to have shaped, cushioned ends
Tubes to be synthetic material

	45
	Stick
	Lightweight aluminum - adjustment range between 75-100cm - PVC handle - anti-slip rubber tip 

	46
	Thermometers 
	Measurement Sites: Oral, axillary, and rectal.
Measurement Time: Approximately 4–6 seconds (oral), 10–13 seconds (rectal/pediatric axillary), 12–15 seconds (adult axillary).
Accuracy: ±0.1 °C (±0.2 °F).
Display: Large, backlit digital LCD with mode indicators and memory recall of last reading.
Probes: Interchangeable color-coded probes (oral/axillary and rectal) with removable probe wells to minimize cross-contamination
Probe Covers: Single-use, disposable, with push-button ejection system.
Battery: Operates on 3 × AA batteries with capacity of approximately 6,000 measurements or rechargeble
Dimensions/Weight: Approx. 215 mm × 81 mm × 62 mm; ~357 g.
Operating Range: 10 °C to 40 °C ambient.
Mounting Options: Handheld

	47
	Trays
	Rectangular tray with rounded corners. Stainless steel. Smooth surface. 30cm*15-22 cm*3 cm. Steel quality, as per ISO standard

	48
	Tripod
	Chrome plated steel - Lightweight approx. 600gr -Height adjustable -  690 to 940mm - Anodized aluminum neck with a molded plastic grip handle - Small base with three footprint approx. 165 X140mm with rubber approx. 3cm on the end -Suitable for left and right handed users ISO certificate

	49
	Tuning forks
	Nickel-plated steel, material thickness 8 mm, for otologists, high-grade steel, nickle-plated
approx. 17 cm

	50
	Ultrasound (cardiac, abdominal, vaginal)
	Color Doppler Ultrasound - Display monitor: minimum 15" color high resolution LCD - Articulating monitor arm - Console design, full size keyboard, 8 TGC pods, Integrated speakers - Probe holders, Gel holder, front and rear handles - Minimum 3 probe connectors - Minimum Built-in storage (HDD) (500GB) - 2 USB ports, DVD-RW driver - Applications: Abdominal, Breast, Obstetrical/Gynecological, cardiac  Real-time automatic doppler calculations 2D imaging - Automatic IMT 
- Probes: Convex transducer, Linear Array ,cardiac adult transducer, Transvaginal - Connectivity: Ethernet port, VGA port - B/W thermal Printer - Electrical power: Voltage 110-220 V AC 
Supplied on stable, mobile trolley fitted with 4 wheels that can be braked
FDA or CE mark.

	51
	Waste Refrigerator
	Cabinet Type: Plastic liner
o   Volume: 550(Liter)
o   Dimensions: (WxDxH) 75x76x180(cm)
o   Number of Shelves: Minimum 3
o   No Freezer
FEATURES:
o   Dynamic cooling system
o   Tropical compressor resist voltage fluctuation.
o   Multi positional shelf system
o   Environmentally friendly
o   Adjustable high load bearing shelves
o   Pre painted steel inner liner
o   Heavy duty unidirectional wheels
o   Adjustable feet in the front
o   Uniform air distribution
o   Wide steel condenser with cooling fan for high efficiency and durability
o   Clean back design
o   Non PVC door frame
o   220/240V 50/60HZ
o   Durable ABS lower grill


	52
	Wheel Chair - pneumatic tires
	Light Weight - Removable wheelchair arms - Soft padded washable upholstery - Detachable swing away wheelchair foot rest - Pneumatic wheelchair tires - Wheels 21’’ + (60cm) in diameter - Locking brakes on each wheel - Detachable wheelchair arms -Folding Back - Overall width:  64cm - Seat width:  45-55cm - Seat depth:  40cm - ISO certificate

	53
	Wheel Chair - Solid tires
	Light Weight - Removable wheelchair arms - Soft padded washable upholstery - Detachable swing away wheelchair foot rest - Solid tires - Wheels 21’’ + (60cm) in diameter - Locking brakes on each wheel -Folding Back - Overall width:  64cm - Seat width:  45-55 cm - Seat depth:  40cm - ISO certificate
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